## General Information
| General Information | I-125 - Bebig/Theragenics model I25.S06
Source Information
Source name: | Bebig/Theragenics model I25.S06
Radionuclide: | I-125
Model Number: | MED3631-A/M
Distributed by: | BEBIG Isotopen und Medizintechnick GmbH, and Theragenics Corporation
Manufactured by:
References
1) | "Update of AAPM Task Group No. 43 Report: A revised AAPM protocol for brachytherapy dose calculations"
Med Phys 31 (3), March 2004, 633-674
M.J. Rivard, B.M. Coursey, L.A. DeWerd, W.F. Hanson, M.S. Huq, G.S. Ibbott, M.G. Mitch, R. Nath,J.F. Williamson
Calculation details
Method: | Monte Carlo and TLD averaged
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised
2006-01-26

## Source design and materials
Geometry and materials | I-125 - Bebig/Theragenics model I25.S06

## Geometry Function
Geometry Function |  |  |  | 125I - Bebig/Theragenics model I25.S06
Active source length,
| L = | 3.5 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | 125I - Bebig/Theragenics model I25.S06
L = | 1.012 | cGy/(h·U)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | 125I - Bebig/Theragenics model I25.S06
| F(r,q) |  | Italicied data are nonconsensus data obtained from candidate datsets
r (cm)
q (º) | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 5 | 7
0 | 0.302 | 0.429 | 0.512 | 0.579 | 0.61 | 0.631 | 0.649 | 0.684
5 | 0.352 | 0.436 | 0.509 | 0.576 | 0.61 | 0.635 | 0.651 | 0.689
10 | 0.44 | 0.476 | 0.557 | 0.622 | 0.651 | 0.672 | 0.689 | 0.721
20 | 0.746 | 0.686 | 0.721 | 0.757 | 0.771 | 0.785 | 0.79 | 0.807
30 | 0.886 | 0.82 | 0.828 | 0.846 | 0.857 | 0.862 | 0.867 | 0.874
40 | 0.943 | 0.897 | 0.898 | 0.907 | 0.908 | 0.913 | 0.918 | 0.912
50 | 0.969 | 0.946 | 0.942 | 0.947 | 0.944 | 0.947 | 0.949 | 0.946
60 | 0.984 | 0.974 | 0.97 | 0.974 | 0.967 | 0.966 | 0.967 | 0.976
70 | 0.994 | 0.989 | 0.988 | 0.99 | 0.984 | 0.985 | 0.987 | 0.994
80 | 0.998 | 0.998 | 0.998 | 1 | 0.994 | 1 | 0.993 | 0.999

## Radial Dose Function
Radial-Dose Function Table |  |  | 125I - Bebig/Theragenics model I25.S06
Line source aproximation,
L = | 3.5 | mm | Point source aproximation,
r (cm) | g(r) |  | r (cm) | g(r)
0.1 | 1.01 |  | 0.1 | 0.613
0.15 | 1.018 |  | 0.15 | 0.76
0.25 | 1.03 |  | 0.25 | 0.908
0.5 | 1.03 |  | 0.5 | 1.001
0.75 | 1.02 |  | 0.75 | 1.012
1 | 1 |  | 1 | 1
1.5 | 0.937 |  | 1.5 | 0.942
2 | 0.857 |  | 2 | 0.863
3 | 0.689 |  | 3 | 0.695
4 | 0.538 |  | 4 | 0.543
5 | 0.409 |  | 5 | 0.413
6 | 0.313 |  | 6 | 0.316
7 | 0.232 |  | 7 | 0.234
8 | 0.176 |  | 8 | 0.178
9 | 0.134 |  | 9 | 0.135
10 | 0.0957 |  | 10 | 0.0967

## Anisotropy Factor
Anisotropy Factor |  |  | 125I - Bebig/Theragenics model I25.S06
Italicied data are nonconsensus data obtained from candidate datsets
| r (cm) | fan(r)
| 0.25 | 1.122
| 0.5 | 0.968
| 1 | 0.939
| 2 | 0.939
| 3 | 0.938
| 4 | 0.94
| 5 | 0.941
| 7 | 0.949

## QA
Dose rates for QA |  |  | I-125 - Bebig/Theragenics model I25.S06
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources using
the 1-D formalism of Eq. (11) with interpolation for gL(r) and Øan(r).
| r (cm) | fan(r)
| 0.5 | 3.922
| 1 | 0.95
| 1.5 | 0.398
| 2 | 0.205
| 3 | 0.0733
| 4 | 0.0323
| 5 | 0.0157
| 6 | 0.0084
| 7 | 0.00459