## General Information
| General Information | I-125 - DraxImage LS-1
Source Information
Source name: | DraxImage LS-1
Radionuclide: | I-125
Model Number: | LS-1 source
Distributed by: | Draximage Inc.
Manufactured by:
References
1) | "Supplement to the 2004 update of the AAPM Task Group Nº43 Report"
Med Phys 34(6), June 2007(2187-2205)
M.J. Rivard, W.M.Butler, L.A. DeWerd, M.S. Huq, G.S Ibbott, A.S. Meigooni, C.S. Melhus, M G. Mitch, R. Nath, J. F. Williamson
Calculation details
Method:
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised

## Geometry and materials
Geometry and materials | I-125 - DraxImage LS-1

## Geometry Function
Geometry Function |  |  | I-125 - DraxImage LS-1
Source length for Line source approximation:
L= | 4.1 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | I-125 - DraxImage LS-1
L = | 0.972 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | I-125 - DraxImage LS-1
F(r,q)
r (cm)
q (º) | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 5 | 10
0 | 3.459 | 1.261 | 0.979 | 0.872 | 0.799 | 0.775 | 0.765 | 0.766 | 0.781
10 | 3.312 | 1.246 | 0.977 | 0.877 | 0.808 | 0.787 | 0.775 | 0.778 | 0.786
20 | 2.755 | 1.219 | 0.988 | 0.901 | 0.841 | 0.821 | 0.811 | 0.816 | 0.822
30 | 2.13 | 1.178 | 0.994 | 0.925 | 0.877 | 0.861 | 0.854 | 0.864 | 0.873
40 | 1.675 | 1.125 | 0.999 | 0.95 | 0.912 | 0.902 | 0.898 | 0.909 | 0.899
50 | 1.38 | 1.073 | 0.998 | 0.967 | 0.945 | 0.938 | 0.934 | 0.94 | 0.935
60 | 1.194 | 1.032 | 0.996 | 0.981 | 0.97 | 0.968 | 0.967 | 0.968 | 0.964
70 | 1.085 | 1.007 | 0.998 | 0.994 | 0.99 | 0.989 | 0.988 | 0.991 | 0.993
80 | 1.024 | 0.999 | 1.001 | 1.001 | 0.999 | 0.999 | 0.998 | 1.004 | 0.982
90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1

## Radial-Dose Function Table
Radial-Dose Function Table |  |  | I-125 - DraxImage LS-1
Line source approximation,
L = | 4.1 | mm | Point source approximation
r (cm) | gL(r) |  | r (cm) | gP(r)
0.1 | 0.182 |  | 0.1 | 0.1
0.15 | 0.323 |  | 0.15 | 0.225
0.25 | 0.741 |  | 0.25 | 0.629
0.5 | 0.964 |  | 0.5 | 0.928
0.75 | 1.004 |  | 0.75 | 0.994
1 | 1 |  | 1 | 1
1.5 | 0.937 |  | 1.5 | 0.944
2 | 0.853 |  | 2 | 0.862
3 | 0.68 |  | 3 | 0.688
4 | 0.527 |  | 4 | 0.534
5 | 0.4 |  | 5 | 0.405
6 | 0.3 |  | 6 | 0.304
7 | 0.223 |  | 7 | 0.226
8 | 0.166 |  | 8 | 0.168
9 | 0.122 |  | 9 | 0.124
10 | 0.09 |  | 10 | 0.0912

## Anisotropy Factor
Anisotropy Factor |  |  | I-125 - DraxImage LS-1
| r (cm) | fan(r)
| 0.1 | 2.004
| 0.15 | 2.275
| 0.25 | 2.152
| 0.5 | 1.15
| 0.75 | 1.03
| 1 | 0.987
| 1.5 | 0.958
| 2 | 0.949
| 3 | 0.943
| 5 | 0.947
| 10 | 0.942

## QA
Dose rates for QA |  |  | I-125 - DraxImage LS-1
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources included in this report using gL(r) and Øan(r),
and the 1-D formalism of Eq. (11) from the 2004 AAPM TG-43U1. Results using interpolated gL(r) or Øan(r) data are highlighted in boldface while
extrapolated results based on gL(r) and/or Øan(r) data are underlined.
| r (cm) | fan(r)
| 0.1 | 19.6
| 0.15 | 22.1
| 0.25 | 21.1
| 0.5 | 4.15
| 0.75 | 1.77
| 1 | 0.959
| 1.5 | 0.39
| 2 | 0.199
| 3 | 0.0701
| 4 | 0.0306
| 5 | 0.0149
| 6 | 0.00781
| 7 | 0.00429
| 8 | 0.00243
| 9 | 0.00141
| 10 | 0.000834