## General Information
| General Information | 125I - Imagyn Medical Technologies Inc. Isostar model IS-12501
Source Information
Source name: | Imagyn Medical Technologies Inc. Isostar model IS-12501
Radionuclide: | I-125
Model Number: | IS-12501
Distributed by: | Imagyn Medical Technologies Inc.
Manufactured by:
References
1) | "Update of AAPM Task Group No. 43 Report: A revised AAPM protocol for brachytherapy dose calculations"
Med Phys 31 (3), March 2004, 633-674
M.J. Rivard, B.M. Coursey, L.A. DeWerd, W.F. Hanson, M.S. Huq, G.S. Ibbott, M.G. Mitch, R. Nath,J.F. Williamson
Calculation details
Method: | Monte Carlo and TLD averaged
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised
2006-01-26

## Source design and materials
Source design and materials | 125I - Imagyn Medical Technologies Inc. Isostar model IS-12501

## Geometry Function
Geometry Function |  |  |  | 125I - Imagyn Medical Technologies Inc. Isostar model IS-12501
Active source length,
| L = | 3.4 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | 125I - Imagyn Medical Technologies Inc. Isostar model IS-12501
L = | 0.94 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | 125I - Imagyn Medical Technologies Inc, Isostar model IS-12501
F(r,q)
r (cm)
q (º) | 1 | 2 | 3 | 5 | 7
0 | 0.241 | 0.337 | 0.362 | 0.424 | 0.454
10 | 0.327 | 0.399 | 0.44 | 0.486 | 0.51
20 | 0.479 | 0.532 | 0.563 | 0.584 | 0.581
30 | 0.634 | 0.663 | 0.681 | 0.706 | 0.7
40 | 0.768 | 0.775 | 0.786 | 0.806 | 0.776
50 | 0.867 | 0.87 | 0.878 | 0.875 | 0.849
60 | 0.946 | 0.944 | 0.944 | 0.943 | 0.913
70 | 0.986 | 0.985 | 0.987 | 0.974 | 0.955
80 | 0.998 | 0.994 | 1.004 | 0.981 | 0.956

## Radial Dose Function
Radial-Dose Function Table |  |  | 125I - Imagyn Medical Technologies Inc, Isostar model IS-12501
Line source aproximation,
L = | 3.4 | mm | Point source aproximation,
r (cm) | g(r) |  | r (cm) | g(r)
0.1 | 1.022 |  | 0.1 | 0.631
0.15 | 1.058 |  | 0.15 | 0.799
0.25 | 1.093 |  | 0.25 | 0.969
0.5 | 1.08 |  | 0.5 | 1.051
0.75 | 1.048 |  | 0.75 | 1.04
1 | 1 |  | 1 | 1
1.5 | 0.907 |  | 1.5 | 0.912
2 | 0.808 |  | 2 | 0.814
3 | 0.618 |  | 3 | 0.623
4 | 0.463 |  | 4 | 0.467
5 | 0.348 |  | 5 | 0.351
6 | 0.253 |  | 6 | 0.255
7 | 0.193 |  | 7 | 0.195
8 | 0.149 |  | 8 | 0.15
9 | 0.1 |  | 9 | 0.101
10 | 0.075 |  | 10 | 0.076

## Anisotropy Factor
Anisotropy Factor |  |  | 125I - Imagyn Medical Technologies Inc, Isostar model IS-12501
| r (cm) | fan(r)
| 1 | 0.867
| 2 | 0.886
| 3 | 0.894
| 5 | 0.897
| 7 | 0.879

## QA
Dose rates for QA |  |  | 125I - Imagyn Medical Technologies Inc. Isostar model IS-12501
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources using the 1-D formalism of Eq. (11) with interpolation for gL(r) and Øan(r). |  |  |  | QA along-away table obtained using TG-43U1S1 formalism and consensus dataset
D(y,z) / cGy h-1 U-1
y / cm
| r (cm) | cGy h-1 U-1 | z / cm | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7
| 0.5 | 3.426 | 0 | --- | 0.94 | 0.19124 | 0.065094 | 0.027445 | 0.013205 | 0.0066675 | 0.0037371
| 1 | 0.815 | 0.5 | 0.89464 | 0.71133 | 0.17586 | 0.062895 | 0.026665 | 0.012842 | 0.0064892 | 0.0036361
| 1.5 | 0.334 | 1 | 0.23551 | 0.36063 | 0.13874 | 0.05558 | 0.024706 | 0.012094 | 0.0061644 | 0.0034716
| 2 | 0.169 | 1.5 | 0.11199 | 0.1741 | 0.097057 | 0.045492 | 0.021922 | 0.011092 | 0.0057897 | 0.003285
| 3 | 0.0582 | 2 | 0.065071 | 0.090665 | 0.064162 | 0.035173 | 0.018518 | 0.009826 | 0.0053284 | 0.0030807
| 4 | 0.0246 | 2.5 | 0.038013 | 0.051115 | 0.042105 | 0.026109 | 0.015088 | 0.0083853 | 0.0047553 | 0.0028407
| 5 | 0.0118 | 3 | 0.023665 | 0.030782 | 0.027843 | 0.019217 | 0.011842 | 0.0069403 | 0.0041339 | 0.0025369
| 6 | 0.00592 | 3.5 | 0.015844 | 0.019475 | 0.018557 | 0.014159 | 0.0092194 | 0.0056191 | 0.0035204 | 0.0022245
| 7 | 0.00328 | 4 | 0.010812 | 0.012794 | 0.012771 | 0.01024 | 0.007043 | 0.0045844 | 0.0029719 | 0.0019141
|  |  | 4.5 | 0.0077739 | 0.0088372 | 0.0088229 | 0.0074267 | 0.0052893 | 0.0036813 | 0.0024876 | 0.0016002
|  |  | 5 | 0.0056075 | 0.0061839 | 0.0061733 | 0.0053513 | 0.0040996 | 0.002925 | 0.0020557 | 0.0013106
|  |  | 5.5 | 0.0040717 | 0.0043996 | 0.0043134 | 0.0039318 | 0.0031461 | 0.0023497 | 0.0016605 | 0.0010483
|  |  | 6 | 0.0029302 | 0.0031567 | 0.0031571 | 0.0029307 | 0.0024365 | 0.0018671 | 0.0013069 | 0.00086055
|  |  | 6.5 | 0.002238 | 0.0023869 | 0.0023561 | 0.0021966 | 0.0019089 | 0.0014595 | 0.0010033 | 0.00071116
|  |  | 7 | 0.001698 | 0.0018044 | 0.0017909 | 0.0016778 | 0.0014789 | 0.0011137 | 0.00078872 | 0.00057945