## General Information
| General Information | I-125 - IsoAid Advantage
Source Information
Source name: | IsoAid Advantage
Radionuclide: | I-125
Model Number: | IAI-125A source
Distributed by:
Manufactured by:
References
1) | "Supplement to the 2004 update of the AAPM Task Group Nº43 Report"
Med Phys 34(6), June 2007(2187-2205)
M.J. Rivard, W.M.Butler, L.A. DeWerd, M.S. Huq, G.S Ibbott, A.S. Meigooni, C.S. Melhus, M G. Mitch, R. Nath, J. F. Williamson
Calculation details
Method:
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised

## Geometry and materials
Geometry and materials | I-125 - IsoAid Advantage

## Geometry Function
Geometry Function |  |  | I-125 - IsoAid Advantage
Source length for Line source approximation:
L = | 3 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | I-125 - IsoAid Advantage
L = | 0.981 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | I-125 - IsoAid Advantage
F(r,q)
r (cm)
q (º) | 0.5 | 1 | 2 | 3 | 5 | 7
0 | 0.352 | 0.406 | 0.493 | 0.52 | 0.578 | 0.612
5 | 0.411 | 0.465 | 0.545 | 0.584 | 0.658 | 0.701
10 | 0.481 | 0.527 | 0.601 | 0.642 | 0.704 | 0.726
20 | 0.699 | 0.719 | 0.757 | 0.775 | 0.794 | 0.799
30 | 0.848 | 0.846 | 0.862 | 0.862 | 0.869 | 0.879
40 | 0.948 | 0.936 | 0.932 | 0.916 | 0.937 | 0.969
50 | 1.002 | 0.986 | 0.974 | 0.961 | 0.963 | 0.971
60 | 1.029 | 1.024 | 1.008 | 0.993 | 0.99 | 1.001
70 | 1.029 | 1.039 | 1.027 | 1.006 | 1.016 | 1.01
80 | 0.999 | 1.025 | 1.024 | 1.023 | 1.009 | 1.025
90 | 1 | 1 | 1 | 1 | 1 | 1

## Radial-Dose Function Table
Radial-Dose Function Table |  |  | I-125 - IsoAid Advantage
Line source approximation,
L = | 3 | mm | Point source approximation
r (cm) | gL(r) |  | r (cm) | gp(r)
0.1 | 1.04 |  | 0.1 | 0.686
0.15 | 1.053 |  | 0.15 | 0.833
0.25 | 1.066 |  | 0.25 | 0.967
0.5 | 1.08 |  | 0.5 | 1.056
0.75 | 1.035 |  | 0.75 | 1.029
1 | 1 |  | 1 | 1
1.5 | 0.902 |  | 1.5 | 0.906
2 | 0.8 |  | 2 | 0.804
3 | 0.611 |  | 3 | 0.615
4 | 0.468 |  | 4 | 0.471
5 | 0.368 |  | 5 | 0.371
6 | 0.294 |  | 6 | 0.296
7 | 0.227 |  | 7 | 0.229
8 | 0.165 |  | 8 | 0.166
9 | 0.141 |  | 9 | 0.142
10 | 0.09 |  | 10 | 0.091

## Anisotropy Factor
Anisotropy Factor |  |  | I-125 - IsoAid Advantage
| r (cm) | fan(r)
| 0.1 | 1.127
| 0.15 | 1.197
| 0.25 | 1.069
| 0.5 | 0.957
| 0.75 | 0.962
| 1 | 0.968
| 2 | 0.964
| 3 | 0.955
| 5 | 0.959
| 7 | 0.955

## QA
Dose rates for QA |  |  | I-125 - IsoAid Advantage
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources included in this report using gL(r) and Øan(r),
and the 1-D formalism of Eq. (11) from the 2004 AAPM TG-43U1. Results using interpolated gL(r) or Øan(r) data are highlighted in boldface while
extrapolated results based on gL(r) and/or Øan(r) data are underlined.
| r (cm) | fan(r)
| 0.1 | 75.9
| 0.15 | 43.5
| 0.25 | 16.2
| 0.5 | 3.97
| 0.75 | 1.73
| 1 | 0.95
| 1.5 | 0.381
| 2 | 0.19
| 3 | 0.064
| 4 | 0.0277
| 5 | 0.0139
| 6 | 0.00772
| 7 | 0.00437
| 8 | 0.00243
| 9 | 0.00164
| 10 | 0.000849