## HDR 12i
| TG-43 CONSENSUS |  |  |  |  | GammaMed HDR 192Ir 12i
Dose calculation for photon-emitting brachytherapy sources with average energy higher than 50 keV: Full Report of the AAPM and ESTRO
Med. Phys. 39 (2012) 2904-2929
| L = | 1.1175 |  | cGy/(h U)
Interpolated / extrapolated data are boldface / underlined. Values inside the source are in italics.
| gL(r) |  |  |  | F(r,q) | [17 empty] | QA Along-away
| L = | 0.35 | cm |  |  |  | Distance from Active Source Center (cm)
| r (cm) | gL(r) L = 0,35 cm unbounded |  | q (º) | 0 | 0.2 | 0.4 | 0.6 | 0.8 | 1 | 1.25 | 1.5 | 1.75 | 2 | 2.5 | 3 | 3.5 | 4 | 5 | 6 | 8 | 10 |  | z (cm) | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
| 0 | 0.99159 |  | 0 | 0.6761 | 0.6761 | 0.6651 | 0.6541 | 0.6392 | 0.633 | 0.6389 | 0.6348 | 0.6459 | 0.637 | 0.6474 | 0.6442 | 0.676 | 0.6769 | 0.7118 | 0.7114 | 0.7537 | 0.7854 |  | 7 | 0.016542 | 0.016855 | 0.017124 | 0.017443 | 0.017694 | 0.01797 | 0.017952 | 0.017073 | 0.015625 | 0.013859 | 0.01212 | 0.010481
| 0.2 | 0.99159 |  | 1 | 0.6648 | 0.6648 | 0.6554 | 0.646 | 0.6347 | 0.631 | 0.6382 | 0.6348 | 0.6464 | 0.6401 | 0.6503 | 0.6476 | 0.6771 | 0.677 | 0.7063 | 0.7068 | 0.7468 | 0.7789 |  | 6 | 0.022128 | 0.022681 | 0.023268 | 0.023801 | 0.024229 | 0.024588 | 0.024332 | 0.022534 | 0.019952 | 0.017212 | 0.014583 | 0.012306
| 0.25 | 0.99164 |  | 2 | 0.7238 | 0.7238 | 0.6936 | 0.6634 | 0.6441 | 0.6404 | 0.6457 | 0.6413 | 0.6524 | 0.6499 | 0.6604 | 0.661 | 0.6897 | 0.693 | 0.719 | 0.7281 | 0.7661 | 0.7904 |  | 5 | 0.032156 | 0.032599 | 0.033432 | 0.034298 | 0.034944 | 0.035289 | 0.034378 | 0.030647 | 0.025922 | 0.021379 | 0.017539 | 0.014362
| 0.5 | 0.99393 |  | 3 | 0.7422 | 0.7422 | 0.7058 | 0.6694 | 0.653 | 0.6506 | 0.6557 | 0.6506 | 0.661 | 0.6612 | 0.6709 | 0.673 | 0.6999 | 0.704 | 0.7241 | 0.734 | 0.7711 | 0.7954 |  | 4 | 0.048087 | 0.050269 | 0.052482 | 0.054049 | 0.054845 | 0.054249 | 0.051201 | 0.042641 | 0.033883 | 0.026552 | 0.020871 | 0.016604
| 0.75 | 0.99681 |  | 4 | 0.7413 | 0.7413 | 0.71 | 0.6787 | 0.6644 | 0.6604 | 0.6638 | 0.6589 | 0.67 | 0.6731 | 0.6823 | 0.6863 | 0.71 | 0.7149 | 0.7312 | 0.744 | 0.7784 | 0.8019 |  | 3 | 0.081692 | 0.087704 | 0.093285 | 0.096122 | 0.096354 | 0.090873 | 0.081508 | 0.060699 | 0.044187 | 0.03252 | 0.024431 | 0.018796
| 1 | 1 |  | 5 | 0.7464 | 0.7464 | 0.7183 | 0.6902 | 0.6769 | 0.6723 | 0.6738 | 0.6678 | 0.6783 | 0.6831 | 0.692 | 0.699 | 0.7212 | 0.727 | 0.7417 | 0.7559 | 0.787 | 0.807 |  | 2 | 0.1822 | 0.2004 | 0.2131 | 0.21361 | 0.20361 | 0.17122 | 0.13683 | 0.085785 | 0.055997 | 0.038535 | 0.027768 | 0.020734
| 1.5 | 1.0042 |  | 6 | 0.769 | 0.769 | 0.7347 | 0.7004 | 0.691 | 0.688 | 0.6884 | 0.681 | 0.691 | 0.6964 | 0.704 | 0.7117 | 0.7308 | 0.7372 | 0.7495 | 0.7653 | 0.7948 | 0.8138 |  | 1.5 | 0.32424 | 0.36087 | 0.38128 | 0.36205 | 0.32537 | 0.24326 | 0.17778 | 0.099938 | 0.061719 | 0.041161 | 0.029139 | 0.021508
| 2 | 1.0058 |  | 7 | 0.7754 | 0.7754 | 0.744 | 0.7126 | 0.7014 | 0.6976 | 0.6987 | 0.6938 | 0.7038 | 0.7099 | 0.7167 | 0.724 | 0.741 | 0.7487 | 0.7602 | 0.7757 | 0.8027 | 0.82 |  | 1 | 0.73711 | 0.85692 | 0.81883 | 0.68673 | 0.54684 | 0.34308 | 0.22468 | 0.11309 | 0.066505 | 0.043266 | 0.03021 | 0.022089
| 3 | 1.0076 |  | 8 | 0.7886 | 0.7886 | 0.755 | 0.7214 | 0.7141 | 0.7132 | 0.7133 | 0.708 | 0.7168 | 0.7228 | 0.7289 | 0.7378 | 0.7527 | 0.7602 | 0.77 | 0.7858 | 0.8106 | 0.8264 |  | 0.5 | 3.3734 | 3.4811 | 2.2127 | 1.3666 | 0.89217 | 0.44965 | 0.26585 | 0.12265 | 0.069734 | 0.044617 | 0.030853 | 0.022441
| 4 | 1.005 |  | 9 | 0.7962 | 0.7962 | 0.7651 | 0.734 | 0.7271 | 0.7253 | 0.7249 | 0.7203 | 0.7286 | 0.734 | 0.7402 | 0.7498 | 0.763 | 0.7709 | 0.781 | 0.796 | 0.8186 | 0.8331 |  | 0 | 389199277.4625 | 15.6259 | 4.317 | 1.9652 | 1.1175 | 0.50156 | 0.28312 | 0.12624 | 0.070859 | 0.045102 | 0.03107 | 0.022562
| 5 | 0.99928 |  | 10 | 0.808 | 0.808 | 0.7763 | 0.7446 | 0.7387 | 0.7366 | 0.7363 | 0.7328 | 0.74 | 0.7461 | 0.751 | 0.7616 | 0.7732 | 0.7814 | 0.7914 | 0.8057 | 0.8264 | 0.8397 |  | -0.5 | 1.6587 | 3.4501 | 2.2024 | 1.3635 | 0.89184 | 0.44908 | 0.26563 | 0.12263 | 0.069646 | 0.044593 | 0.030815 | 0.022427
| 6 | 0.99116 |  | 15 | 0.8521 | 0.8521 | 0.8292 | 0.8063 | 0.801 | 0.7969 | 0.7948 | 0.7963 | 0.8 | 0.8026 | 0.8067 | 0.8147 | 0.8216 | 0.829 | 0.8391 | 0.8487 | 0.8621 | 0.8712 |  | -1 | 0.39528 | 0.83743 | 0.81282 | 0.68219 | 0.54438 | 0.34194 | 0.22413 | 0.11288 | 0.066355 | 0.04321 | 0.030138 | 0.022062
| 8 | 0.96804 |  | 20 | 0.9737 | 0.9737 | 0.8761 | 0.855 | 0.849 | 0.8447 | 0.8433 | 0.8466 | 0.8492 | 0.8484 | 0.853 | 0.8561 | 0.8622 | 0.8681 | 0.8751 | 0.8811 | 0.8899 | 0.897 |  | -1.5 | 0.18111 | 0.34472 | 0.37621 | 0.35898 | 0.3237 | 0.24209 | 0.17704 | 0.099722 | 0.061531 | 0.041063 | 0.029062 | 0.021474
| 10 | 0.93627 |  | 30 | 0.981 | 0.981 | 0.9308 | 0.9204 | 0.917 | 0.914 | 0.9139 | 0.9108 | 0.9138 | 0.9126 | 0.9146 | 0.9147 | 0.9184 | 0.9202 | 0.9222 | 0.9274 | 0.931 | 0.9358 |  | -2 | 0.10575 | 0.18525 | 0.20848 | 0.21193 | 0.2024 | 0.17044 | 0.13626 | 0.085558 | 0.055803 | 0.038417 | 0.027689 | 0.020697
|  |  |  | 40 | 0.9921 | 0.9921 | 0.9648 | 0.9553 | 0.9539 | 0.9519 | 0.9513 | 0.9489 | 0.9499 | 0.95 | 0.9494 | 0.9504 | 0.9537 | 0.9539 | 0.9546 | 0.9578 | 0.9591 | 0.962 |  | -3 | 0.050718 | 0.076972 | 0.088957 | 0.09384 | 0.09509 | 0.090267 | 0.080966 | 0.060312 | 0.043971 | 0.032439 | 0.024365 | 0.018751
|  |  |  | 50 | 0.9949 | 0.9949 | 0.9797 | 0.9742 | 0.9746 | 0.974 | 0.9724 | 0.9717 | 0.973 | 0.9737 | 0.9718 | 0.9726 | 0.9733 | 0.9732 | 0.9751 | 0.9756 | 0.9766 | 0.9784 |  | -4 | 0.030562 | 0.042554 | 0.048752 | 0.051884 | 0.053444 | 0.053451 | 0.050653 | 0.042362 | 0.033756 | 0.026486 | 0.0208 | 0.016542
|  |  |  | 60 | 0.9979 | 0.9979 | 0.9897 | 0.9879 | 0.9879 | 0.987 | 0.985 | 0.9837 | 0.9877 | 0.9867 | 0.9859 | 0.986 | 0.9867 | 0.9867 | 0.988 | 0.9881 | 0.9885 | 0.9889 |  | -5 | 0.020801 | 0.027476 | 0.030805 | 0.03274 | 0.033812 | 0.034596 | 0.033889 | 0.030429 | 0.025803 | 0.021291 | 0.017456 | 0.014301
|  |  |  | 70 | 0.9989 | 0.9989 | 0.9952 | 0.995 | 0.9946 | 0.9948 | 0.9942 | 0.992 | 0.9954 | 0.9944 | 0.9941 | 0.9942 | 0.9956 | 0.9953 | 0.9947 | 0.9962 | 0.9952 | 0.995 |  | -6 | 0.015263 | 0.018808 | 0.021146 | 0.02255 | 0.023371 | 0.024026 | 0.023919 | 0.022305 | 0.019821 | 0.017117 | 0.014505 | 0.012247
|  |  |  | 80 | 1.0017 | 1.0017 | 1 | 0.9999 | 0.9992 | 0.999 | 0.999 | 0.9983 | 0.998 | 0.9984 | 0.9984 | 0.9983 | 0.9997 | 0.9995 | 0.9988 | 1.0003 | 0.9993 | 0.9998 |  | -7 | 0.011764 | 0.013834 | 0.015387 | 0.016344 | 0.016951 | 0.017493 | 0.017609 | 0.016851 | 0.015502 | 0.013768 | 0.012048 | 0.010425
|  |  |  | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  |  | 100 | 0.9997 | 0.9997 | 0.9989 | 0.9983 | 0.9984 | 0.9992 | 0.9983 | 0.997 | 0.9988 | 0.9982 | 0.998 | 0.9983 | 0.9974 | 0.9977 | 0.9979 | 0.9977 | 0.9988 | 0.998
|  |  |  | 110 | 0.9979 | 0.9979 | 0.9947 | 0.9936 | 0.993 | 0.9948 | 0.993 | 0.9912 | 0.9928 | 0.9926 | 0.9923 | 0.9929 | 0.992 | 0.9922 | 0.9919 | 0.993 | 0.9936 | 0.9938
|  |  |  | 120 | 0.9966 | 0.9966 | 0.989 | 0.9868 | 0.9863 | 0.9863 | 0.9852 | 0.9822 | 0.9843 | 0.9838 | 0.9833 | 0.9842 | 0.985 | 0.9836 | 0.9837 | 0.986 | 0.9847 | 0.986
|  |  |  | 130 | 0.9969 | 0.9969 | 0.9777 | 0.9708 | 0.971 | 0.9693 | 0.969 | 0.967 | 0.9703 | 0.9683 | 0.9671 | 0.968 | 0.9701 | 0.968 | 0.9701 | 0.9738 | 0.9717 | 0.974
|  |  |  | 140 | 0.9931 | 0.9931 | 0.9598 | 0.949 | 0.9493 | 0.9471 | 0.9459 | 0.945 | 0.9471 | 0.946 | 0.9453 | 0.9473 | 0.9476 | 0.9463 | 0.9489 | 0.9542 | 0.9537 | 0.956
|  |  |  | 150 | 0.9778 | 0.9778 | 0.9244 | 0.9119 | 0.909 | 0.9086 | 0.9059 | 0.9047 | 0.9076 | 0.9087 | 0.9098 | 0.9117 | 0.9129 | 0.912 | 0.915 | 0.9206 | 0.9238 | 0.9294
|  |  |  | 160 | 0.9448 | 0.9448 | 0.8912 | 0.8411 | 0.838 | 0.8392 | 0.836 | 0.8369 | 0.8388 | 0.842 | 0.8448 | 0.8467 | 0.8543 | 0.8553 | 0.8611 | 0.8676 | 0.8759 | 0.8883
|  |  |  | 165 | 0.8532 | 0.8532 | 0.817 | 0.7808 | 0.7818 | 0.7826 | 0.7819 | 0.7829 | 0.786 | 0.7884 | 0.7946 | 0.7983 | 0.8069 | 0.811 | 0.8207 | 0.8311 | 0.8453 | 0.859
|  |  |  | 170 | 0.828 | 0.828 | 0.765 | 0.702 | 0.713 | 0.7026 | 0.7067 | 0.704 | 0.7127 | 0.7148 | 0.7243 | 0.73 | 0.74 | 0.7476 | 0.7623 | 0.7792 | 0.8012 | 0.8173
|  |  |  | 171 | 0.775 | 0.775 | 0.738 | 0.701 | 0.689 | 0.69 | 0.687 | 0.6844 | 0.6937 | 0.6962 | 0.705 | 0.712 | 0.7232 | 0.7313 | 0.7478 | 0.7659 | 0.7897 | 0.8069
|  |  |  | 172 | 0.772 | 0.772 | 0.724 | 0.676 | 0.676 | 0.678 | 0.6689 | 0.6643 | 0.674 | 0.6766 | 0.6838 | 0.6914 | 0.7048 | 0.7134 | 0.732 | 0.7511 | 0.7768 | 0.7952
|  |  |  | 173 | 0.657 | 0.657 | 0.654 | 0.651 | 0.648 | 0.657 | 0.627 | 0.6416 | 0.6521 | 0.655 | 0.6611 | 0.6694 | 0.6851 | 0.6941 | 0.7146 | 0.7339 | 0.7619 | 0.7828
|  |  |  | 174 | 0.635 | 0.635 | 0.63 | 0.625 | 0.62 | 0.618 | 0.601 | 0.62 | 0.6284 | 0.6326 | 0.6363 | 0.6454 | 0.663 | 0.6726 | 0.695 | 0.7146 | 0.7449 | 0.7688
|  |  |  | 175 | 0.671 | 0.671 | 0.647 | 0.623 | 0.599 | 0.575 | 0.581 | 0.601 | 0.6 | 0.6071 | 0.6085 | 0.618 | 0.638 | 0.6466 | 0.6722 | 0.6903 | 0.7233 | 0.7518
|  |  |  | 176 | 0.616 | 0.616 | 0.602 | 0.588 | 0.574 | 0.56 | 0.544 | 0.569 | 0.588 | 0.556 | 0.5758 | 0.5876 | 0.6104 | 0.6191 | 0.648 | 0.6652 | 0.7001 | 0.7334
|  |  |  | 177 | 0.2424 | 0.2424 | 0.2808 | 0.3192 | 0.3576 | 0.396 | 0.444 | 0.492 | 0.553 | 0.5 | 0.527 | 0.54 | 0.5659 | 0.5767 | 0.6139 | 0.6337 | 0.6713 | 0.7113
|  |  |  | 178 | 0.9586 | 0.9586 | 0.9002 | 0.8418 | 0.7834 | 0.725 | 0.652 | 0.579 | 0.506 | 0.433 | 0.448 | 0.479 | 0.55 | 0.54 | 0.583 | 0.588 | 0.635 | 0.6816
|  |  |  | 179 | 0.4948 | 0.4948 | 0.4956 | 0.4964 | 0.4972 | 0.498 | 0.499 | 0.5 | 0.501 | 0.502 | 0.504 | 0.506 | 0.508 | 0.51 | 0.522 | 0.5719 | 0.5842 | 0.6416
|  |  |  | 180 | 0.31525 | 0.31525 | 0.3213 | 0.32735 | 0.3334 | 0.33945 | 0.34701 | 0.35457 | 0.36214 | 0.3697 | 0.38482 | 0.39995 | 0.41507 | 0.4302 | 0.46045 | 0.4907 | 0.5512 | 0.6117