## General Information
| General Information | Pd-103 - Best Medical model 2335 source
Source Information
Source name: | Best Medical model 2335 source
Radionuclide: | Pd-103
Model Number: | 2335 source
Distributed by:
Manufactured by:
References
1) | "Supplement to the 2004 update of the AAPM Task Group Nº43 Report"
Med Phys 34(6), June 2007(2187-2205)
M.J. Rivard, W.M.Butler, L.A. DeWerd, M.S. Huq, G.S Ibbott, A.S. Meigooni, C.S. Melhus, M G. Mitch, R. Nath, J. F. Williamson
Calculation details
Method:
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised

## Geometry and materials
Geometry and materials | Pd-103 - Best Medical model 2335 source

## Geometry Function
Geometry Function |  |  | Pd-103 - Best Medical model 2335 source
Source length for Line source approximation:
L= | 4.55 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | Pd-103 - Best Medical model 2335 source
L = | 0.685 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | Pd-103 - Best Medical model 2335 source
F(r,q)
r (cm)
q (º) | 1 | 2 | 3 | 4 | 5 | 6 | 7
0 | 0.797 | 0.69 | 0.674 | 0.672 | 0.663 | 0.675 | 0.63
5 | 0.801 | 0.696 | 0.683 | 0.669 | 0.666 | 0.679 | 0.645
10 | 0.79 | 0.69 | 0.673 | 0.675 | 0.665 | 0.69 | 0.644
15 | 0.675 | 0.613 | 0.608 | 0.604 | 0.626 | 0.62 | 0.581
20 | 0.608 | 0.591 | 0.596 | 0.601 | 0.616 | 0.647 | 0.595
25 | 0.675 | 0.639 | 0.637 | 0.659 | 0.653 | 0.706 | 0.651
30 | 0.681 | 0.66 | 0.679 | 0.694 | 0.694 | 0.717 | 0.672
35 | 0.725 | 0.693 | 0.705 | 0.721 | 0.703 | 0.73 | 0.687
40 | 0.762 | 0.736 | 0.75 | 0.747 | 0.741 | 0.775 | 0.72
45 | 0.792 | 0.807 | 0.846 | 0.847 | 0.866 | 0.876 | 0.804
50 | 0.885 | 0.88 | 0.885 | 0.887 | 0.909 | 0.907 | 0.835
60 | 0.915 | 0.929 | 0.944 | 0.936 | 0.965 | 1.001 | 0.912
70 | 0.932 | 0.96 | 0.972 | 0.965 | 0.975 | 1.014 | 0.916
80 | 0.941 | 0.975 | 0.986 | 0.985 | 0.999 | 1.017 | 0.915
90 | 1 | 1 | 1 | 1 | 1 | 1 | 1

## Radial-Dose Function Table
Radial-Dose Function Table |  |  | Pd-103 - Best Medical model 2335 source
Line source approximation,
L = | 4.55 | mm | Point source approximation
r (cm) | gL(r) |  | r (cm) | gp(r)
0.1 | 0.826 |  | 0.1 | 0.427
0.15 | 1.066 |  | 0.15 | 0.706
0.25 | 1.236 |  | 0.25 | 1.02
0.5 | 1.307 |  | 0.5 | 1.247
0.75 | 1.128 |  | 0.75 | 1.114
1 | 1 |  | 1 | 1
1.5 | 0.742 |  | 1.5 | 0.749
2 | 0.533 |  | 2 | 0.539
3 | 0.296 |  | 3 | 0.3
4 | 0.158 |  | 4 | 0.161
5 | 0.092 |  | 5 | 0.0935
6 | 0.0529 |  | 6 | 0.0538
7 | 0.0309 |  | 7 | 0.0314
8 | 0.018 |  | 8 | 0.0184
9 | 0.0105 |  | 9 | 0.0107
10 | 0.0062 |  | 10 | 0.0063

## Anisotropy Factor
Anisotropy Factor |  |  | Pd-103 - Best Medical model 2335 source
| r (cm) | fan(r)
| 0.1 | 1.052
| 0.15 | 1.205
| 0.25 | 1.213
| 0.5 | 0.938
| 0.75 | 0.894
| 1 | 0.879
| 2 | 0.872
| 3 | 0.881
| 4 | 0.881
| 5 | 0.89
| 6 | 0.909
| 7 | 0.845

## QA
Dose rates for QA |  |  | Pd-103 - Best Medical model 2335 source
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources included in this report using gL(r) and Øan(r),
and the 1-D formalism of Eq. (11) from the 2004 AAPM TG-43U1. Results using interpolated gL(r) or Øan(r) data are highlighted in boldface while
extrapolated results based on gL(r) and/or Øan(r) data are underlined.
| r (cm) | fan(r)
| 0.1 | 30.8
| 0.15 | 25.9
| 0.25 | 12.6
| 0.5 | 3.21
| 0.75 | 1.21
| 1 | 0.602
| 1.5 | 0.2
| 2 | 0.0806
| 3 | 0.0201
| 4 | 0.00605
| 5 | 0.00228
| 6 | 0.00093
| 7 | 0.000371
| 8 | 0.000166
| 9 | 7.66e-05
| 10 | 3.62e-05